## Summary
FCMPASS Report Summary
FCMPASS Version: | 4.2
Dataset Date: | 2023-01-09
Dataset Name: | FluoSpheres Calibration
Number of Files Calibrated: | 2
Number of Output Figures: | 1

## MIFlowCyt
Framework Criteria | What to report | Please complete each criterion
1.1. Purpose | A brief description of the goal of the experiment. This should include the rationale and hypothesis. | Following FCMPASS protocols allows for the determination of optimal FL and scatter gain for analyzing small particles. The lower limit of detection can also be quantified with proper controls detailed in resources 1 and 2.
1.2. Keywords | The keywords should include terms from an appropriate vocabulary (e.g., MeSH) to describe the experiment. | N/A
1.3. Experiment Variables | Variables are attribute(s) that differ between samples within an experiment due to pre-existing differences in sample states or due to experimental manipulation of the samples. For additional information, consult the supplementary information of the references [1]. | N/A, all samples transferred from the same stock diluted using commercially available reagents.
1.4. Organization name and address | List the experimental primary contact name and email. | NIH/NCI, 10 Center Drive, Bethesda, MD, 20892
1.5. Primary contact name and email address | The description may include information for additional individuals involved in the experiment, including their contact details and their role. | Joshua Welsh, joshua.welsh@nih.gov
1.6. Date or time period of experiment | The date or time period during which the investigation was performed (i.e., from collecting and treating samples to performing data analysis) shall be stated. | 09-Jan-2023
1.7. Conclusions | A brief summary of the interpretation of the results or outcome of the experiment shall be provided if available. | Scatter and fluorescent optimal gains determined for CytoFLEX and Aurora on violet side scatter and FITC channels by determining greatest separation of beads from noise population.
1.8. Quality control measures | A description of the quality control measures used such as replicates, calibrations, control assays, etc. shall be provided. If another experiment was performed for the purposes of establishing quality control standards, that experiment may be referenced. | See MiFlowCytEV sections 4.3 and 4.4.
2.1.1.1. (2.1.2.1., 2.1.3.1.) Sample description | Sample materials shall be described according to 2.1.1, 2.1.2, or 2.1.3, based on the type of the sample material. Each sample shall be distinguished from other samples within the same experiment. For additional information, consult the supplementary information of the references [1]. | 0.1 µm FluoSpheres (https://www.thermofisher.com/order/catalog/product/F8807)
2.1.1.2. Biological sample source description | The source of the biological sample shall be described (e.g., wildtype mouse, C57BL/6 spleen, C57BL/6 splenocytes). If the source is a cell line the description shall include its name, ATCC [17] (or equivalent) number, and cell type. | N/A
2.1.1.3. Biological sample source organism description | See supplementary information of reference and consult sections 2.1.1.3.1-2.1.1.3.7 for detailed descriptions of expected information [1]. | N/A
2.1.2.2. Environmental sample location | The location of the sample origin shall be specified. This should include longitude and latitude if relevant. | N/A
2.3. Sample treatment description | The description shall include details about treatment agents, which play the role of experiment variables, or shall specify that samples were untreated. | N/A
2.4. Fluorescence reagent(s) description | The expected and possible types of cells or other particles in the sample material, and their expected and possible measurable characteristics should be noted as these represent the key information for choosing appropriate reagents with respect to the experimental goal. The following information shall be provided about each fluorescence reagent used. Information about the characteristic(s) being measured, as well as details about what plays the role of analyte, analyte detector, and analyte reporter should be descriptive enough to allow for accurate interpretation of the experiment data. | N/A
3.1. Instrument manufacturer | The instrument manufacturer shall be specified by stating its name (e.g., MIFlowCyt Cytometry Systems, Inc.). A URL pointing to manufacturer web pages may also be provided. | Beckman Coulter
3.2. Instrument model | The instrument model shall be specified (e.g., CytMeter Elite-2000). It may include the instrument serial number. | CytoFLEX S
3.3. Instrument configuration and settings | The system and its configurations have significant impact on experimental results and as such the following shall be provided. Please consult the supplementary information of the reference for additional information [1]. | Refer to Sample Acq Info and Instr Settings sheets for a list of data files.
4.1. List-mode data files: We recommend all authors to submit their data files to http://flowrepository.org and to make them available for the peer-review process | The list-mode data files (e.g., FCS files) shall be provided directly or details on how they may be requested shall be stated. | Refer to MIFlowCytEV section 7.1
4.2. Compensation description | A description of the type of compensation used shall be included (e.g., no compensation, hardware compensation, computed compensation) and the spillover or compensation matrix shall be provided when possible (i.e., it may not be available for old data but shall be provided when available). While the spillover matrix is preferred, the compensation matrix is also acceptable. The type of the matrix (i.e., spillover vs. compensation) shall be explicitly stated. | N/A
4.3. Data transformation details | Describe each data/parameter transformation performed during analysis when the transformation does not qualify as compensation. | Refer to MIFlowCytEV sections 4 and 5 for detailed description of data transformation.
4.4.1. Gate description | The subpopulation identified by the gate shall be briefly described and the gating strategy or a reference to where it is described in detail should be provided | Refer to gating scripts or .wsp files for gating methods.
4.4.2. Gate statistics | Percentage of events within the gate shall be provided specifically stating the denominator. The denominator shall be either the total population of events or another gate. When the denominator is another gate, this shall also be exactly specified in recursive fashion so that the gating strategy can be followed up to the original data set. | Refer to gating scripts or .wsp files for gating methods.
4.4.3. Gate boundaries | Either the exact mathematical descriptions of each gate boundary shall be provided using an open and freely available specification, or this information shall be provided in the form of images. | Refer to gating scripts or .wsp files for gating methods.
References
Lee, J. A. et al. MIFlowCyt: the minimum information about a Flow Cytometry Experiment. Cytometry. Part A : the journal of the International Society for Analytical Cytology, 73(10), 926–930. https://doi.org/10.1002/cyto.a.20623

## MIFlowCytEV
Framework Criteria | What to report | Please complete each criterion
1.1 Preanalytical variables conforming to MISEV guidelines. | Preanalytical variables relating to EV sample including source, collection, isolation, storage, and any others relevant and available in the performed study. | N/A
1.2 Experimental design according to MIFlowCyt guidelines. | EV-FC manuscripts should provide a brief description of the experimental aim, keywords, and variables for the performed FC experiment(s) using MIFlowCyt checklist criteria: 1.1, 1.2, and 1.3, respectively. Template found at www.evflowcytometry.org. | See 'MIFlowCyt' sheet.
2.1 Sample staining details | State any steps relating to the staining of samples. Along with the method used for staining, provide relevant reagent descriptions as listed in MIFlowCyt guidelines (Section 2.4 Fluorescence Reagent(s) Descriptions). | N/A, all samples diluted from commercially available reagents and no further staining was performed.
2.2 Sample washing details | State any steps relating to the washing of samples. | N/A
2.3 Sample dilution details | All methods and steps relating to sample dilution. | DPBS acquired at same settings.
3.1 Buffer alone controls. | State whether a buffer-only control was analyzed at the same settings and during the same experiment as the samples of interest. If utilized it is recommended that all samples be recorded for a consistent set period of time e.g. 5 minutes, rather than stopping analysis at a set recorded event count e.g. 100,000 events. This allows comparisons of total particle counts between controls and samples. | N/A, no staining performed
3.2 Buffer with reagent controls. | State whether a buffer with reagent control was analyzed at the same settings, same concentrations, and during the same experiment as the samples of interest. If used state what the results were. | N/A, no staining performed.
3.3 Unstained controls. | State whether unstained control samples were analyzed at the same settings and during the same experiment as stained samples. If used, state what the results were, preferably in standard units. | N/A, no staining performed.
3.4 Isotype controls. | The use of isotype controls is applicable to immunofluorescence labelling only. State whether isotype controls were analyzed at the same settings and during the same experiment as stained samples. If utilized, state which antibody they are matched to, the concentration used, and what the results were (Section 4.2, 4.3, 4.4). Due to conjugation differences between manufacturers if should be stated if the isotype controls are from the same manufacturer as the matched antibodies. | N/A, no staining performed.
3.5 Single-stained controls. | State whether single-stained controls were included. If used state whether the single-stained controls were recorded using the same settings, dilutions, and during the same experiment as stained samples and state what the results were, preferably in standard units (Section 4.2, 4.3, 4.4). | N/A, no staining performed.
3.6 Procedural controls. | State whether procedural controls were included. If used, state the procedure and if the procedural controls were acquired at the same settings and during the same experiment as stained samples. | N/A, no staining performed.
3.7 Serial dilutions. | State whether serial dilutions were performed on samples and note the dilution range and manner of testing. The fluorescence and/or scatter signal intensity would ideally be reported in standard units (see Section 4.3, 4.4) but arbitrary units can also be used. This data is best reported by plotting the recorded number events/concentration over a set period of time at different sample dilution. The median fluorescence intensity at each of the dilutions should also ideally be plotted on the same or a separate plot. | N/A
3.8. Detergent treated EV-samples | State whether samples were detergent treated to assess lability. If utilized, state what detergent was used, the end concentration of the detergent, and what the results were of the lysis. | N/A
4.1 Trigger Channel(s) and Threshold(s). | The trigger channel(s) and threshold(s) used for event detection. Preferably, the fluorescence calibration (Section 4.3) and/or scatter calibration (Section 4.4) should be used in order to report the trigger channel(s) and threshold(s) in standardized units. | Trigger Channel(s) and Threshold(s) in calibrated and arbitrary units are summarized for each sample in the table on the 'Sample Acq Info' sheet.
4.2 Flow Rate / Volumetric quantification. | State if the flow rate was quantified/validated and if so, report the result and how they were obtained. | low flow rate
4.3 Fluorescence Calibration. | State whether fluorescence calibration was implemented, and if so, report the materials and methods used, catalogue numbers, lot numbers, and supplied reference units for the standards. Fluorescence parameters may be reported in standardized units of MESF, ERF, or ABC beads. The type of regression used, and the resulting scatter plot of arbitrary data vs standard data for the reference particles should be supplied. | Fluorescence calibration was not performed.
4.4 Light Scatter Calibration. | State whether and how light scatter calibration was implemented. Light scatter parameters may be reported in standardized units of nm2, along with information required to reproduce the model. | Light scatter calibration was performed using FCMPASS software as previously demonstrated [1,2,3]. Details on the reagents and modelling parameters can be found in the summary table on the 'SSC Cali' sheet and in FCMPASS Figure 1.
5.1 EV diameter/surface area/volume approximation. | State whether and how EV diameter, surface area, and/or volume has been calculated using FC measurements. | Diameter approximation using light scattering was performed using FCMPASS software as previously demonstrated [1,2,3]. Details of modelling parameters for calibration, core-shell, and solid sphere models can be found in the 'SSC Cali' sheet and in FCMPASS Figure 1.
5.2 EV refractive index approximation. | State whether the EV refractive index has been approximated and how this was done. | N/A
5.3 EV epitope number approximation. | State whether EV epitope number has been approximated, and if so, how it was approximated. | N/A
6.1 Completion of MIFlowCyt checklist. | Complete MIFlowCyt checklist criteria 1 to 4 using the MIFlowCyt guidelines. Template found at www.evflowcytometry.org. | See 'MIFlowCyt' sheet.
6.2 Calibrated channel detection range | If fluorescence or scatter calibration has been carried out, authors should state whether the upper and lower limits of a calibrated detection channel were calculated in standardized units. This can be done by converting the arbitrary unit scale to a calibrated scaled, as discussed in Section 4.3 and 4.4, and providing the highest unit on this scale and the lowest detectable unit above the unstained population. The lowest unit at which a population is deemed ‘positive’ can be determined a variety of ways, including reporting the 99th percentile measurement unit of the unstained population for fluorescence. The chosen method for determining at what unit an event was deemed positive should be clearly outlined. | N/A
6.3 EV number/concentration. | State whether EV number/concentration has been reported. If calculated, it is preferable to report EV number/concentration in a standardized manner, stating the number/concentration between a set detection range. | N/A
6.4 EV brightness. | When applicable, state the method by which the brightness of EVs is reported in standardized units of scatter and/or fluorescence. | Refer to 'SSC Cali' sheet.
7.1. Sharing of data to a public repository. | Provide a link to the experimental data in a public data repository. | repo link scrub
References
Welsh, J.A., Horak, P., Wilkinson, J.S., Ford, V.J., Jones, J.C., Smith, D., Holloway, J.A. and Englyst, N.A. (2019), FCMPASS Software Aids Extracellular Vesicle Light Scatter Standardization. Cytometry. doi:10.1002/cyto.a.23782
Welsh, J.A., Jones, J.C. and Tang, V.A. (2020), Fluorescence and Light Scatter Calibration Allow Comparisons of Small Particle Data in Standard Units across Different Flow Cytometry Platforms and Detector Settings. Cytometry. doi:10.1002/cyto.a.24029
Welsh, J. A., & Jones, J. C. (2020). Small particle fluorescence and light scatter calibration using FCMPASS software. Current Protocols in Cytometry, 94, e79. doi: 10.1002/cpcy.79

## Sample Acq Info
The following extrapolations of thresholds to standard units assume that samples were acquired at the same detector settings as their calibration controls.
If applied, calibration metadata for fluorescence and light scattering calibration can be found in the 'Fl Cali' and 'SSC Cali' sheets, respectively.
Filename | Sample Type | Sample Description | Trigger Parameter | Trigger ID | Trigger Threshold (au) | Detector Setting (au)
FluoSpheres Ref8803Lot1437608.fcs | Sample |  | Violet SSC | SSC_1-H | 1500 | 250
NIST Bead Mix.fcs | Sample |  | Violet SSC | SSC_1-H | 1500 | 250

## SSC Cali
Flow Cytometer Modelling Settings
SSC Parameter | Wavelength | Sheath RI | Determine HA | Aperture Geom | Theta | Phi | Eps | Cali. Factor (nm^2)
SSC_1-A ¦ Violet SSC-A | 405 | 1.3431 | off | circle | 90 | 90 | 55.2 | 0.0052515
Light scatter calibration reference beads
Diameter (nm) | Diameter CV (%) | Measured RI | Measurement Wavelength | Composition | Acquired Stat | Acquired CV | Input | Acquisition Wavelength | Modelled RI | Manufacturer | Catalogue No. | Lot No.
100 | 6.8 | 1.59 | 589 | Polystyrene | 16602 |  | Manual | 405 | 1.6253 | ThermoFisher Scientific | 3100A | 204935
152 | 3.3 | 1.59 | 589 | Polystyrene | 111328 |  | Manual | 405 | 1.6253 | ThermoFisher Scientific | 3150A | 202026
240 | 7.8 | 1.59 | 589 | Polystyrene | 544922 |  | Manual | 405 | 1.6253 | ThermoFisher Scientific | 3200A | 205131
303 | 1.6 | 1.59 | 589 | Polystyrene | 931641 |  | Manual | 405 | 1.6253 | ThermoFisher Scientific | 3300A | 204665
401 | 1.3 | 1.59 | 589 | Polystyrene | 1148438 |  | Manual | 405 | 1.6253 | ThermoFisher Scientific | 3400A | 203859
Core-Shell Modelling Settings
Shell RI | Core RI | Shell Thickness (nm) | Parameter Name
1.4863 | 1.406 | 5 | EV Diameter (nm) [High RI]
1.4863 | 1.3859 | 5 | EV Diameter (nm) [Average RI]
1.4863 | 1.3658 | 5 | EV Diameter (nm) [Low RI]
1.4863 | 1.42 | 5 | rEV CS Diameter (nm)
References
Welsh, J.A., Horak, P., Wilkinson, J.S., Ford, V.J., Jones, J.C., Smith, D., Holloway, J.A. and Englyst, N.A. (2019), FCMPASS Software Aids Extracellular Vesicle Light Scatter Standardization. Cytometry. doi:10.1002/cyto.a.23782
Welsh, J.A., Jones, J.C. and Tang, V.A. (2020), Fluorescence and Light Scatter Calibration Allow Comparisons of Small Particle Data in Standard Units across Different Flow Cytometry Platforms and Detector Settings. Cytometry. doi:10.1002/cyto.a.24029
Welsh, J. A., & Jones, J. C. (2020). Small particle fluorescence and light scatter calibration using FCMPASS software. Current Protocols in Cytometry, 94, e79. doi: 10.1002/cpcy.79

## Instr Settings
Instrument Settings
Attribute Name | Attribute Info
Serial Number | BE02004
Files | Parameter | Gain (au)
| FSC-H | FSC-A | FSC-Width | SSC-H | SSC-A | SSC_1-H | SSC_1-A | FL1-H | FL1-A | FL2-H | FL2-A | FL3-H | FL3-A | FL4-H | FL4-A | FL5-H | FL5-A | FL6-H | FL6-A | FL7-H | FL7-A | FL8-H | FL8-A | FL9-H | FL9-A | FL10-H | FL10-A | FL11-H | FL11-A | FL12-H | FL12-A | Time
FluoSpheres Ref8803Lot1437608.fcs | 100 | 100 | Not Specified | 100 | 100 | 250 | 250 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | Not Specified
NIST Bead Mix.fcs | 100 | 100 | Not Specified | 100 | 100 | 250 | 250 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | 500 | Not Specified